## Start Here
EU MDR Checklist
Last refreshed: 03 March 2026
Get a complete overview of the key changes, timelines, and important steps
to become fully compliant with EU MDR. With this checklist,
you can undergo a self-assessment to check your EU MDR readiness.
MDR transition timing has changed since this checklist was first issued.
Regulation (EU) 2023/607 extended the transitional period for eligible 'legacy devices' (Article 120) beyond 26 May 2024, with risk‑based end dates (see below).
Key end dates (if Article 120 conditions are met): • 31 Dec 2027 – Class III and most implantable Class IIb devices (excluding certain WET implants) • 31 Dec 2028 – other Class IIb, Class IIa, Is/Im and certain up‑classified devices
To benefit from the extension, manufacturers must (among other things) have an MDR‑aligned QMS in place by 26 May 2024, lodge an MDR conformity assessment application by 26 May 2024, and sign a written agreement with an MDR Notified Body by 26 Sept 2024. The MDR 'sell‑off' deadline has been removed (subject to shelf‑life/expiry).
EUDAMED roll‑out update: the first four EUDAMED modules become mandatory from 28 May 2026 (Actor registration, UDI/device registration, Notified Bodies & Certificates, Market Surveillance).
The remaining modules (Post‑Market Surveillance & Vigilance; Clinical Investigations/Performance Studies) will become mandatory when released (no voluntary period announced).
1. Check the Sheet named "EU MDR Checklist - Scilife" of this document.
2. Go through the 14 different requirements and its corresponding EU MDR Sections.
3. Check the first cell if your medical device meets the requirements.
You are all set! | Click here to know more
UPDATE LINK Don't miss any important detail by checking the entire blog post here

## EU MDR Change log
Date | What changed | Primary reference
2023-03 | Regulation (EU) 2023/607 entered into force, extending MDR Article 120 transitional provisions for eligible legacy devices and removing the MDR 'sell‑off' deadline. Review eligibility, deadlines, and evidence needed. | EC Article 120 flowchart (Aug 2023) + EC Q&A on Reg (EU) 2023/607
2023-12 | MDCG 2019‑7 Rev.1 (PRRC guidance) published. Confirm PRRC appointment, qualifications, scope, and interfaces (manufacturer + AR). | MDCG 2019‑7 Rev.1
2024-06 | Regulation (EU) 2024/1860 introduced a 'gradual roll‑out' mechanism for EUDAMED modules becoming mandatory once independently audited and declared functional. | Regulation (EU) 2024/1860
2024-10 | MDCG 2021‑25 Rev.1 (legacy devices guidance) updated. Check applicability of MDR requirements to legacy devices placed on the market during the transition. | MDCG 2021‑25 Rev.1
2025-11 | Commission notice/decision declaring the functionality of the first four EUDAMED modules; starts the 6‑month transition period leading to mandatory use from 28 May 2026. | European Commission EUDAMED overview page
2025-12 | Commission issued a proposal to amend/simplify aspects of MDR/IVDR (watchlist only until adopted). | COM(2025) 1023 (proposal)

## EU MDR Checklist - Scilife
www.scilife.io
Last refreshed: 03 March 2026
Requirements / Details |  | EU MDR requirements description | Related EU MDR section | Comments & Additional Sources
1 | Risk Management System | Implement a risk management system to conduct the risk management process throughout the life cycle of the device, including design and process risk assessment, updated (as appropriate) with data from PMS (post-market surveillance). The assessment should show that benefits outweigh residual risks when the device is used as intended. Required documents: 1. Risk management procedure(s), including the definition of any rating systems 2. RMP (risk management plan) 3. RMF (risk management file) 4. RMR (risk management report) 5. Risk analysis (e.g., hazard analysis, FMEA), including appropriate controls (i.e., protective measures) for all identified risks | Annex I Chapter I | ISO 14971:2019 (and ISO/TR 24971:2020 guidance)
2 | PMS (Post-Market Surveillance) & Vigilance | Implement an MDR‑compliant Post‑Market Surveillance (PMS) and Vigilance system. Typical deliverables (risk‑class dependent): 1. PMS procedure(s) + PMS Plan (Annex III) 2. PMS Report (Art. 85) and/or PSUR (Art. 86) 3. Trend reporting procedure (Art. 88) 4. Vigilance reporting process (serious incidents, FSCAs/FSNs) incl. timelines (Arts. 87–92) 5. PMCF Plan + PMCF Evaluation Report (Annex XIV Part B) 6. Signal management + periodic review / management review inputs Tip: align PMS outputs with Clinical Evaluation updates and risk management reviews. | Chapter VII Section I (Arts. 83–86) Chapter VII Section II (Arts. 87–92) Chapter VII Section III (Arts. 93–95, 97); Annex III; Annex XIV Part B | ISO 13485:2016 (PMS/Vigilance processes)
3 | Economic Operators | Document and implement responsibilities, agreements and interfaces for all applicable economic operators. Include (as applicable): • Manufacturer obligations (Art. 10) incl. QMS, technical documentation, PMS, vigilance, UDI, registration • Person Responsible for Regulatory Compliance (PRRC) (Art. 15) • Authorised Representative mandate and tasks (Art. 11) • Importer and distributor obligations (Arts. 13–14) • Verification steps (CE mark, DoC, UDI, labelling/IFU, storage/transport) and complaint handling • Traceability: keep records of suppliers/customers as required Deliverables often include: role‑specific SOPs, AR mandate, quality agreements, and operator checklists. | Chapter II (Arts. 10–16); Annex VI Part A; Article 31 | MDCG 2019‑7 Rev.1 (PRRC guidance); ISO 13485:2016 (outsourced processes & supplier control)
4 | Clinical Evaluation and Clinical Investigations | Maintain MDR‑compliant clinical evaluation (CER) and, where applicable, clinical investigations. Typical deliverables: 1. Clinical Evaluation Plan (CEP) 2. Clinical Evaluation Report (CER) (Annex XIV Part A) 3. Clinical development strategy / clinical evidence mapping 4. PMCF Plan + PMCF Evaluation Report (Annex XIV Part B) 5. For Class III and implantable devices: Summary of Safety and Clinical Performance (SSCP) (Art. 32) 6. Clinical Investigation documentation (where applicable) (Annex XV) Tip: keep CER, PMS/PMCF and risk management tightly linked and consistently updated. | Article 61; Annex XIV (Part A & B); Article 32; Annex XV | MEDDEV 2.7/1 rev.4 (legacy reference); MDCG clinical evaluation guidance (see MDCG endorsed docs list)
5 | EU DoC (Declaration of Conformity) | Ensure the EU Declaration of Conformity (DoC) is complete and kept up‑to‑date. Must include (per Annex IV), e.g.: • Manufacturer + SRN (if available) • Basic UDI‑DI (where applicable) • Device identification + risk class + conformity assessment route • Notified Body details + certificates (if applicable) • Reference to relevant common specifications / harmonised standards used | Annex IV Article 10(8)
6 | CE (EU Compliance Mark) | Where compliance with applicable requirements has been demonstrated following the applicable conformity assessment procedure, manufacturers of devices other than custom-made or investigational devices shall clearly affix the CE mark to the device of conformity in accordance with Article 20. | Annex V Chapter II Articles 10(6) and 20(1–6)
7 | Classification Rules | The classification route and its rationale must address each requirement of the decided classification rule. Definitions specific to classification rules: 1. Duration of use 2. Invasive and active devices Implementing rules: 1. Implementing rules Classification Rules: 1. Non-Invasive devices 2. Invasive devices 3. Active devices 4. Special rules | Annex VIII Chapters I–III
8 | Routes of Conformity | Manufacturers of Class I devices, other than custom-made or investigational devices, shall declare the conformity of their products by issuing the EU declaration of conformity referred to in Article 19 after drawing up the technical documentation set out in Annexes II and III. If those devices are placed on the market in sterile condition, have a measuring function or are reusable surgical instruments, the manufacturer shall apply the procedures set out in Chapters I and III of Annex IX, or in Part A of Annex XI. | Annex XI Section 2 Article 52
9 | Quality Management System | The quality management system shall address at least the following aspects that are stated in Article 10, Paragraph 9, items from a to m. (a) a strategy for regulatory compliance, including compliance with conformity assessment procedures and procedures for management of modifications to the devices covered by the system (b) identification of applicable general safety and performance requirements and exploration of options to address those requirements (c) responsibility of the management (d) resource management, including selection and control of suppliers and sub-contractors (e) risk management as set out in in Section 3 of Annex I (f) clinical evaluation in accordance with Article 61 and Annex XIV, including PMCF (g) product realisation, including planning, design, development, production, and service provision (h) verification of the UDI assignments made in accordance with Article 27(3) to all relevant devices and ensuring consistency and validity of information provided in accordance with Article 29 (i) setting-up, implementation and maintenance of a post-market surveillance system, in accordance with Article 83 (j) handling communication with competent authorities, notified bodies, other economic operators, customers and/or other stakeholders (k) processes for reporting of serious incidents and field safety corrective actions in the context of vigilance (l) management of corrective and preventive actions and verification of their effectiveness (m) processes for monitoring and measurement of output, data analysis and product improvement. | Annex IX Chapter I Annex XI Part A Clause 6 Article 10(9) | ISO 13485:2016 (baseline); align with Article 10(9) MDR requirements
10 | GSPRs (General Safety and Performance Requirements) | General requirements The technical file should include GSPRs evaluation of Annex I to ensure and show the conformity. A checklist that includes all requirements can be used for this requirement. Required documents 1. Safety & performance statement | Annex I and Annex II
|  | Requirements regarding design and manufacture - Chemical, physical, and biological properties - Infection and microbial contamination - Devices incorporating a substance considered to be a medicinal product and devices that are composed of substances or of combinations of substances that are absorbed by or locally dispersed in the human body. - Devices incorporating materials of biological origin - Construction of devices and interaction with their environment - Devices with a diagnostic or measuring function - Protection against radiation General - Electronic programmable systems — devices that incorporate electronic programmable systems and software that are devices in themselves - Active devices and devices connected to them - Requirements for active implantable devices - Protection against mechanical and thermal risks - Protection against the risks posed to the patient or user by devices supplying energy or substances - Protection against the risks posed by medical devices intended by the manufacturer for use by laypersons | Annex I Chapter II Clauses 10–22
|  | The Information supplied with the device - Information on the label - Information on the packaging which also maintain the sterile condition of a device (if available) - Information in the instructions for use - If an electronic IFU will be provided, organization should ensure the compliance with all relevant aspects of Regulation 207/2012. Required documents - Each type of labelling (primary, secondary, tertiary packaging) including all required symbols. - Sterile information on the packaging - IFU (Instructions for use) Documents may also include following - surgical manuals, brochures, etc. | Annex I Chapter III Clause 23 Annex II Clause 1 | ISO 15223-1:2021 Medical devices — Symbols to be used with information to be supplied by the manufacturer — Part 1: General requirements Regulation (EU) 207/2012 - electronic instructions for use of medical devices
11 | UDI | Implement UDI assignment, labelling and traceability requirements. Deliverables (as applicable): 1. UDI procedure(s) including UDI‑DI / UDI‑PI assignment rules 2. Label/IFU updates to include UDI carrier(s) and Basic UDI‑DI (where required) 3. Direct marking plan for reusable devices (if applicable) 4. UDI data submissions / device registration preparations (EUDAMED UDI/Device module) Note: UDI obligations can apply during the extended transitional period for legacy devices depending on the scenario. | Article 10(9)(h); Articles 27–29; Annex VI; Annex XII Chapter I; Article 27(7) | MDR Article 27; Annex VI Part C; Annex XII; see EC Q&A on 2023/607 (UDI during transition)
12 | EUDAMED | Prepare for the gradual roll‑out of EUDAMED obligations. From 28 May 2026, mandatory use of the first four modules: • Actor registration • UDI/Device registration • Notified Bodies & Certificates • Market Surveillance Key actions: 1. Obtain/verify SRN(s) and actor data; ensure operator roles are correct 2. Prepare UDI/Device master data for registration (including Basic UDI‑DI/UDI‑DI relationships) 3. Align certificate management and vigilance/PMS reporting plans with EUDAMED rollout The remaining modules (PMS & Vigilance; Clinical Investigations/Performance Studies) will become mandatory when released. | Articles 26, 33–34; EUDAMED gradual roll-out per Regulation (EU) 2024/1860 and Commission Decision (EU) 2025/2371 | European Commission EUDAMED overview + timeline (module mandatory dates)
13 | Technical Documentation | Device description and specification, Including all variants, configurations, and accessories Required documents 1.Device description and specification (a) product or trade name and a general description (b) the Basic UDI-DI (c) the intended patient population and medical conditions to be diagnosed, treated and/or monitored and other considerations (d) principles of operation of the device and its mode of action (e) the rationale for the qualification of the product as a device (f) the risk class of the device and the justification for the classification rule(s) (g) an explanation of any novel features (h) a description of the accessories for a device (i) a description or complete list of the various configurations/variants (j) a general description of the key functional elements (k) a description of the raw materials incorporated into key functional elements and those making either direct contact with the human body or indirect contact with the body, (l) technical specifications, 2.Reference to previous and similar generations of the device (a) an overview of the previous generation or generations of the device (b) an overview of identified similar devices available on the Union or international markets | Annex II Clause 1
|  | Information to be supplied by the manufacturer A complete set of: — the label or labels on the device and on its packaging, such as single unit packaging, sales packaging, transport packaging in case of specific management conditions, in the languages accepted in the Member States where the device is envisaged to be sold; and Required documents - Labels IFU | Annex II Clause 2
|  | Design and manufacturing information Complete information and specifications, including the manufacturing processes and their validation, their adjuvants, the continuous monitoring, and the final product testing. Data shall be fully included in the technical documentation Required documents - Design information Manufacturing process and validations - Site identification | Annex II Clause 3
|  | Product Verification and Validation 1. Pre-clinical and clinical data 2. Additional information required in specific cases Required documents - Verification and validation of pre-clinical and clinical testing | Annex II Clause 6
14 | MDR transition / 'legacy device' planning | Confirm whether your device is a 'legacy device' eligible for extended transition under Regulation (EU) 2023/607 (Article 120) and plan MDR certification accordingly. Key dates (if Article 120 conditions are met): • Until 31 Dec 2027 – Class III and most implantable Class IIb devices (excluding certain WET implants) • Until 31 Dec 2028 – other Class IIb, Class IIa, Is/Im and certain up‑classified devices Core conditions to benefit from the extension (non‑exhaustive): • Maintain compliance with the Directives (MDD/AIMDD) and avoid significant changes in design/intended purpose • Device must not present an unacceptable risk to health or safety • Put in place an MDR‑aligned QMS by 26 May 2024 • Lodge an MDR conformity assessment application by 26 May 2024 • Sign a written agreement with an MDR Notified Body by 26 Sept 2024 Note: the MDR 'sell‑off' deadline was removed – devices lawfully placed on the market can continue to be made available/put into service (subject to shelf‑life/expiry). | Article 120(3a–3f); Regulation (EU) 2023/607 | EC Article 120 flowchart (Aug 2023) + EC Q&A on Reg (EU) 2023/607